## Sheet1
This file contains power calibration data for microscopy light sources, collected via a Thorlabs PM100D Optical Power Meter.
Notes | SPECTRAX data are taken across the Cy5 filter cube mirror (T660lpxr), or the Cherry2 mirror (FF605-Di02) | [22 empty] | Original SPECTRAX Calibration Data
[24 empty] | Cy5 Mirror Data | [5 empty] | Cherry2 Mirror Data
Mfr. | Lumencor | Lumencor | Lumencor | Lumencor | Lumencor | Lumencor | Lumencor | Lumencor | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | Lumencor (4251) | [6 empty] | Mfr.
Model | SOLA | SOLA | SOLA | SOLA | SOLA | SOLA | SOLA | SOLA | SPECTRAX_Blue | SPECTRAX_Teal | SPECTRAX_Violet | SPECTRAX_Blue | SPECTRAX_Cyan | SPECTRAX_Teal | SPECTRAX_Green | SPECTRAX_Red | [6 empty] | Model | Violet | Blue | Cyan | Teal | Green | Red | Violet | Blue | Cyan | Teal | Violet | Cal Wl Correction
Filter | DAPI | CFP | GFP | YFP | Orange | Cherry3 | Cy5 | Tsap | CFPa | YFP | NONE | NONE | NONE | NONE | NONE | NONE | [6 empty] | Filter | NONE | NONE | NONE | NONE | NONE | NONE | NONE | NONE | NONE | NONE | NONE
Ex. Line | NONE | NONE | NONE | NONE | NONE | NONE | NONE | NONE | 2 | 4 | 1 | 2 | 3 | 4 | 5 | 6 | [6 empty] | Ex. Line | 1 | 2 | 3 | 4 | 5 | 6 | 1 | 2 | 3 | 4 | 1
Cal. WL | 395 | 435 | 480 | 500 | 545 | 586 | 635 | 405 | 435 | 514 | 395 | 435 | 480 | 514 | 545 | 635 | Old calibrations | [5 empty] | Cal. WL | 395 | 435 | 480 | 514 | 545 | 635 | 395 | 435 | 480 | 514 | 435 (oops) |  |  | Teal Cy5 Mirror
Obj. | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x |  |  |  |  | SPECTRAX |  | Obj. | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x | 20x |  |  | vs.
Pwr % | mW | mW | mW | mW | mW | mW | mW | mW | mW | mW | [10 empty] | Green ND4 |  | Pwr % | mW | mW | mW | mW | mW | mW | mW | mW | mW | mW | mW |  |  | Teal Cherry2 Mirror | bar: 1 to 6
1 | 0.32 | 0.21 | 0.3 | 0.12 | 0 | 0.0001717 | 0 | 0.3468 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | From light guide directly |  | Via YFP cube | YFP scaled | 0 | #DIV/0! | 1 | 0 | 0 | 0.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.155 | Avg. | 1.0342
2 | 0.97 | 1.36 | 1.34 | 0.49 | 0.03 | 0.01306 | 0.26 | 1.008 | 0.46 | 0.01 | 0.53752 | 0.66186 | 1.4375 | 0.17581 | 0.45 | 0.33 | Power % | mW |  |  | 0.12 | 3.75 | 2 | 0.3 | 0.42 | 1.16 | 0.23 | 0.45 | 0.33 | 0.51977 | 0.64 | 1.39 | 0.17 | 0.45 |  |  | 1.3529
3 | 1.3 | 2.01 | 1.66 | 0.76 | 0.06 | 0.02917 | 0.48 | 1.353 | 1.2 | 0.03 | 0.98245 | 1.6443 | 1.8511 | 0.45503 | 1.04 | 0.56 | 1 | 5.5 |  |  | 0.26 | 4 | 3 | 0.55 | 1 | 1.5 | 0.49 | 1.04 | 0.56 | 0.95 | 1.59 | 1.79 | 0.44 | 0.83 | 1.1446 |  | 1.1136
4 | 1.76 | 2.73 | 2.15 | 0.96 | 0.15 | 0.08141 | 0.82 | 1.827 | 1.91 | 0.05 | 1.5648 | 2.5957 | 2.4923 | 0.7963 | 2.09 | 0.92 | 2 | 31.1 |  |  | 0.53 | 3.9434 | 4 | 0.86 | 1.59 | 2.01 | 0.83 | 2.09 | 0.92 | 1.5131 | 2.51 | 2.41 | 0.77 | 1.31 |  |  | 1.0779
5 | 2.07 | 3.2 | 2.51 | 1.1 | 0.26 | 0.162 | 1.01 | 2.138 | 2.37 | 0.07 | 1.9132 | 3.2162 | 2.968 | 1.0342 | 2.94 | 1.15 | 3 | 43.8 |  |  | 0.74 | 3.973 | 5 | 1.07 | 1.97 | 2.38 | 1.05 | 2.94 | 1.15 | 1.85 | 3.11 | 2.87 | 1 | 1.62 | 1.142 |  | 1.05
7 | 2.84 | 4.38 | 3.42 | 1.46 | 0.6 | 0.3725 | 1.45 | 2.934 | 3.53 | 0.12 | 2.8548 | 4.8192 | 4.1263 | 1.6133 | 6.06 | 1.65 | 4 | 62.5 |  |  | 1.53 | 3.9608 | 7 | 1.57 | 2.96 | 3.3 | 1.6 | 6.06 | 1.65 | 2.7606 | 4.66 | 3.99 | 1.56 | 2.39 |  |  | 1.0256
10 | 4.03 | 6.06 | 4.86 | 2 | 1.42 | 0.8523 | 2.16 | 4.149 | 5.19 | 0.22 | 4.2194 | 7.1564 | 5.9257 | 2.4923 | 11.2 | 2.42 | 5 | 75.7 | 0.557 | 75.7 | 2.8 | 4 | 10 | 2.35 | 4.41 | 4.73 | 2.39 | 11.2 | 2.42 | 4.08 | 6.92 | 5.73 | 2.41 | 3.57 | 1.1429 |  | 0.9917
12 | 4.76 | 7.08 | 5.76 | 2.32 | 1.96 | 1.192 | 2.59 | 4.886 | 6.18 | 0.28 | 5.1483 | 8.5835 | 7.0219 | 2.9163 | 14.3 | 2.91 | 7 | 111 |  |  | 3.58 | 3.9944 | 12 | 2.83 | 5.31 | 5.6 | 2.88 | 14.3 | 2.91 | 4.9782 | 8.3 | 6.79 | 2.82 | 4.31 |  |  | 1.0213
15 | 5.92 | 8.64 | 7.15 | 2.8 | 2.81 | 1.746 | 3.31 | 6.058 | 7.66 | 0.39 | 6.4945 | 10.7552 | 8.7386 | 3.7023 | 19.2 | 3.68 | 10 | 167 | 1.02 | 138.6248 | 4.8 | 4 | 15 | 3.6 | 6.67 | 6.94 | 3.57 | 19.2 | 3.68 | 6.28 | 10.4 | 8.45 | 3.58 | 5.28 | 1.1894 |  | 0.99721
20 | 7.59 | 10.8 | 9.17 | 3.47 | 4.07 | 2.572 | 4.36 | 7.761 | 9.72 | 0.55 | 8.5884 | 13.7543 | 11.1689 | 4.6227 | 26.3 | 4.84 | 12 | 204 |  |  | 6.58 | 3.997 | 20 | 4.72 | 8.63 | 8.82 | 4.5 | 26.3 | 4.84 | 8.3048 | 13.3 | 10.8 | 4.47 | 7.19 |  |  | 1.0067
25 | 9.34 | 13.1 | 11.3 | 4.15 | 5.43 | 3.442 | 5.5 | 9.536 | 11.8 | 0.71 | 10.6518 | 16.9602 | 13.6509 | 5.5431 | 33.4 | 6.09 | 15 | 263 | 1.46 | 198.4237 | 8.44 | 3.9573 | 25 | 5.89 | 10.7 | 10.8 | 5.4 | 33.4 | 6.09 | 10.3 | 16.4 | 13.2 | 5.36 | 8.97 | 1.1483 |  | 1.0075
30 | 11.1 | 15.2 | 13.3 | 4.76 | 6.72 | 4.289 | 6.62 | 11.23 | 13.7 | 0.88 | 12.7811 | 19.9592 | 15.926 | 6.3601 | 40.5 | 7.33 | 20 | 352 |  |  | 10.3 | 3.932 | 30 | 7.01 | 12.6 | 12.6 | 6.2 | 40.5 | 7.33 | 12.359 | 19.3 | 15.4 | 6.15 | 10.7 |  |  | 1.0081
35 | 12.7 | 17.2 | 15.2 | 5.33 | 7.99 | 5.109 | 7.72 | 12.84 | 15.4 | 1.04 | 14.6923 | 22.7514 | 18.2012 | 7.0943 | 47 | 8.53 | 25 | 443 | 2.17 | 294.9174 | 12 | 3.9167 | 35 | 8.1 | 14.5 | 14.3 | 6.93 | 47 | 8.53 | 14.2071 | 22 | 17.6 | 6.86 | 12.3 |  |  | 1.0102
40 | 14.1 | 18.9 | 16.9 | 5.83 | 9.14 | 5.852 | 8.73 | 14.33 | 16.9 | 1.18 | 16.4841 | 25.2334 | 20.1661 | 7.6631 | 54.4 | 9.64 | 30 | 533 |  |  | 13.6 |  | 40 | 9.08 | 16.1 | 15.8 | 7.55 | 0 | 9.64 | 15.9396 | 24.4 | 19.5 | 7.41 | 13.8 |  |  | 1.0189
45 | 15.6 | 20.7 | 18.7 | 6.34 | 10.4 | 6.634 | 9.8 | 15.86 | 18.5 | 1.32 | 18.2758 | 27.7154 | 22.131 | 8.3146 | 61.2 | 10.8 | 35 | 618 | 2.8 | 380.5386 | 15.3 |  | 45 | 10.1 | 17.8 | 17.3 | 8.17 | 0 | 10.8 | 17.6722 | 26.8 | 21.4 | 8.04 | 15.3 |  |  | 1.0162
50 | 17.1 | 22.4 | 20.4 | 6.81 | 11.5 | 7.395 | 10.9 | 17.35 | 19.9 | 1.46 | 20.0626 | 30.1974 | 24.0959 | 8.9144 | 67.6 | 11.9 | 40 | 696 |  |  | 16.9 |  | 50 | 11.1 | 19.4 | 18.7 | 8.74 | 4.619 | 11.9 | 19.4 | 29.2 | 23.3 | 8.62 | 16.8 | 1.1548 |  | 1.0139
55 | 18.5 | 23.9 | 22 | 7.25 | 12.7 | 8.151 | 11.9 | 18.77 | 21.4 | 1.59 | 21.7399 | 32.5759 | 25.8539 | 9.4625 | 73.6 | 13 | 45 | 776 |  |  | 18.4 |  | 55 | 12.1 | 21 | 20.1 | 9.28 | 0 | 13 | 21.0218 | 31.5 | 25 | 9.15 | 18.2 |  |  | 1.0142
60 | 19.8 | 25.3 | 23.4 | 7.64 | 13.7 | 8.831 | 12.8 | 20.01 | 22.5 | 1.7 | 23.2927 | 34.6443 | 27.4052 | 9.9486 | 79.2 | 14 | 50 | 853 | 3.59 | 487.9048 | 19.8 |  | 60 | 12.9 | 22.4 | 21.3 | 9.76 | 0 | 14 | 22.5234 | 33.5 | 26.5 | 9.62 | 19.5 |  |  | 1.0146
65 | 21.1 | 26.7 | 24.8 | 8.05 | 14.8 | 9.544 | 13.8 | 21.34 | 23.8 | 1.82 | 24.8455 | 36.816 | 29.0598 | 10.445 | 85.2 | 15.2 | 55 | 928 |  |  | 21.3 |  | 65 | 13.8 | 23.9 | 22.5 | 10.3 | 0 | 15.2 | 24.0249 | 35.6 | 28.1 | 10.1 | 20.8 |  |  | 1.0198
70 | 22.4 | 28.1 | 26.3 | 8.42 | 15.8 | 10.23 | 14.9 | 22.62 | 25 | 1.94 | 26.3984 | 38.8843 | 30.611 | 10.8586 | 90.8 | 16.2 | 60 | 995 |  |  | 22.7 |  | 70 | 14.7 | 25.3 | 23.6 | 10.7 | 0 | 16.2 | 25.5265 | 37.6 | 29.6 | 10.5 | 22.1 |  |  | 1.019
75 | 23.6 | 29.4 | 27.6 | 8.79 | 16.8 | 10.93 | 15.8 | 23.84 | 26.1 | 2.05 | 27.9222 | 40.8492 | 32.0589 | 11.3757 | 96 | 17.1 | 65 | 1070 |  |  | 24 |  | 75 | 15.5 | 26.7 | 24.7 | 11.1 | 4.6352 | 17.1 | 27 | 39.5 | 31 | 11 | 23.3 | 1.1588 |  | 1.0091
80 | 24.7 | 30.5 | 28.8 | 9.11 | 17.8 | 11.53 | 16.5 | 24.88 | 27.1 | 2.14 | 29.1457 | 42.6073 | 33.2998 | 11.7894 | 100.8 | 18 | 70 | 1140 |  |  | 25.2 |  | 80 | 16.3 | 27.9 | 25.6 | 11.5 | 0 | 18 | 28.1831 | 41.2 | 32.2 | 11.4 | 24.4 |  |  | 1.0088
85 | 25.8 | 31.7 | 30 | 9.45 | 18.7 | 12.18 | 17.4 | 26.09 | 28.1 | 2.24 | 30.5791 | 44.3653 | 34.6443 | 12.203 | 106 | 18.8 | 75 | 1210 | 4.66 | 633.325 | 26.5 |  | 85 | 17.1 | 29.1 | 26.6 | 11.9 | 0 | 18.8 | 29.5691 | 42.9 | 33.5 | 11.8 | 25.6 |  |  | 1.0085
90 | 26.9 | 32.8 | 31.2 | 9.77 | 19.7 | 12.82 | 18.2 | 27.15 | 29.1 | 2.34 | 31.8931 | 46.1234 | 35.8852 | 12.5133 | 111.2 | 19.7 | 80 | 1270 |  |  | 27.8 |  | 90 | 17.8 | 30.4 | 27.5 | 12.3 | 0 | 19.7 | 30.8397 | 44.6 | 34.7 | 12.1 | 26.7 |  |  | 1.0165
95 | 27.9 | 33.9 | 32.3 | 10.1 | 20.7 | 13.42 | 18.9 | 28.17 | 30 | 2.43 | 33.207 | 47.778 | 37.0228 | 12.927 | 116 | 20.4 | 85 | 1330 |  |  | 29 |  | 95 | 18.6 | 31.6 | 28.4 | 12.7 | 0 | 20.4 | 32.1102 | 46.2 | 35.8 | 12.5 | 27.8 |  |  | 1.016
100 | 28.9 | 34.8 | 33.3 | 10.4 | 21.4 | 13.92 | 19.6 | 29.06 | 30.9 | 2.52 | 34.5408 | 49.2259 | 38.057 | 13.2372 | 119.6 | 21.1 | 90 | 1390 |  |  | 29.9 |  | 100 | 19.3 | 32.6 | 29.2 | 13 | 4.6389 | 21.1 | 33.4 | 47.6 | 36.8 | 12.8 | 28.8 | 1.1597 |  | 1.0156
[17 empty] | 95 | 1450
[17 empty] | 100 | 1500 | 5.51 | 748.8456